## Slide 1: Photoacoustic and Ultrasonic Tomography for Breast Imaging
Felix Lucka*, Ben Cox &
the PAMMOTH team (*he/him, felix.lucka@cwi.nl)
Medical Physics Seminar @ MLU Halle-Wittenberg, 16th Nov 2023

## Slide 2: Breast Cancer
Most common cause of cancer death in women worldwide:
25% of all cancer cases
15% of all cancer deaths
Early detection and diagnosis are key to successful treatment!

## Slide 3: Conventional Breast Imaging
Magnetic resonance imaging
contrast agent
expensive
availability
not suitable for every woman
Ultrasonography
whole breast imaging is difficult
operator dependent
courtesy of Rianne Bulthuis and Maura Dantuma
X-ray mammography
painful compression
ionizing radiation (tumor induction)
glandular tissue is problematic
Urgent need for novel imaging techniques providing higher specificity, contrast and image resolution than X-ray mammography at lower costs than MRI.

## Slide 4: Photoacoustic Tomography (PAT)
adapted from Rianne Bulthuis and Maura Dantuma
Hybrid imaging: light in, sound out
non-ionizing, near-infrared radiation
quantitative images of optical properties
novel diagnostic information

## Slide 5: Photoacoustic Imaging: Spectral Properties
Oxyhemoglobin
Deoxyhemoglobin
Elastin
Collagen
Lipid (a) (b)
Water
(from Beard, 2011)
different wavelengths allow quantitative spectroscopic examinations
gap between oxygenated and deoxygenated blood
use of contrast agents for molecular imaging

## Slide 6: Ultrasound Tomography (UST)
adapted from Rianne Bulthuis and Maura Dantuma
sound in, sound out
different from conventional US but as safe
quantitative images of acoustic properties
novel diagnostic information

## Slide 7: Speed of Sound vs MRI Images
From: Duric, Littrup, 2017. Breast Ultrasound Tomography, IntechOpen

## Slide 8: H2020: PAMMOTH
Combined PAT+UST scanner to obtain novel diagnostic information from high resolution maps of optical and acoustic properties

## Slide 9: PAM3 System

## Slide 10: Our Contributions
simulation studies for
ultrasonic transducer specification
light excitation design
sensing pattern design
measurement protocol design
reconstruction algorithm design
accuracy vs. computational time/resources/complexity
scanner modelling
assist high performance computing implementation
assist phantom & calibration design
process data, refine measurement procedures

## Slide 11: PAT: Mathematical Modelling
radiative transfer equation (RTE)
photoacoustic effect
acoustic wave equation
(assuming thermal and stress confinement conditions)
measurement (on boundary)
two coupled inverse problems:
acoustic initial value problem with boundary data.
optical parameter identification with internal data.

## Slide 12: PAT Image Reconstruction – General Considerations
Imaging goal - what do we want?
qualitative (structural) or quantitative images?
fast (real-time?) image formation?
Data properties - where and how did we measure the data?
sensing geometry – planar, cylindrical, (half-)spherical
spatio-temporal sampling rates, compression?
sensor response, noise level
Wave propagation model - how realistic do we need to be?
need to include sound speed / density variations?
need to account for absorption?
include shear waves?
What else?
available computing hardware?
pre- and postprocessing

## Slide 13: PAT Sensing Geometries
taken from: Paul Beard Interface Focus 2011;1:602-631
What is their difference?

## Slide 14: PAT Visibility Condition
taken from: Xu, Y. et al. (2004). Med. Phys., 31(4), 724.
visibility region
A phase space point is “visible” (“audible”), if a ray through in the direction of hits a sensor within the measurement time.
“Visibility region”: all points such that is visible for all .

## Slide 15: PAT Reconstruction Approaches
Ad-hoc back-projection / delay-and-sum / beamforming
early approaches translated from US imaging
simple geometric space-time correspondence
efficient & fast to implement
Analytical methods / series solutions
based on rigorous mathematics
exact for certain geometries and conditions
sometimes fast implementations (FFTs)
Model-based (algebraic / iterative / variational / Bayesian)
flexible, minimal conditions, sparse / compressed data
explicit numerical forward model, a-priori information
computationally heavy (high-res 3D?)
Learned / data-driven (machine/deep learning)
improve certain aspect of traditional approaches
use training data to learn good inversion / image representation

## Slide 16: Model-Based Approaches
models the wave propagation
models the sensing geometry and characteristics
combined forward model
Idea: Solve linear system approximately and subject to additional constraints
cannot be set up as explicit matrix, but matrix-vector products , can be computed  iterative solution (e.g. conjugate gradient)
, require numerical wave propagation model
Arridge, Betcke, Cox, L, Treeby, 2016. On the Adjoint Operator in Photoacoustic Tomography. Inverse Problems

## Slide 17: Acoustic Wave Solvers
Direct methods
finite-difference
pseudospectral
finite/spectral element
discontinous Galerkin
Integral equation methods
boundary element
Asymptotic methods
geometrical optics
Gaussian beams
Time domain vs. frequency domain

## Slide 18: k-Wave
Time-domain k-space pseudospectral method solving the underlying system of first order conservation laws
compute spatial derivatives via 3D FFTs
parallel/GPU computing
modify finite temporal differences by k-space operator + staggered grids for accuracy and robustness
perfectly matched layer  free-space propagation
Treeby, Cox, 2010. k-Wave: MATLAB toolbox for the simulation and reconstruction of photoacoustic wave fields. Journal of Biomedical Optics

## Slide 19: Variational Regularization Approaches
Idea: Formulate image reconstruction as variational energy-minimization to incorporate a-priori information on :
hard constraints encode physical information, e.g.,
(weighted) least-squares data-fit encodes Gaussian noise
penalizes un-wanted features of
can be solved by first-order optimization methods
to compute gradient, ones needs matrix-vector products ,
Chambolle, Pock, 2016. An introduction to continuous optimization for imaging. Acta Numerica

## Slide 20: Comparison for One-Sided Sensor Data

## Slide 21: Comparison for Sub-Sampled Data

## Slide 22: PAMMOTH modeling, baseline
Notes: SOS model: cup

## Slide 23
PAMMOTH modeling, + spatial response
Notes: SOS model: cup

## Slide 24
PAMMOTH modeling, + temporal response
Notes: SOS model: cup

## Slide 25
PAMMOTH modeling, + sampling density
Notes: SOS model: cup

## Slide 26
PAMMOTH modeling, higher resolution
Notes: SOS model: cup

## Slide 27: Illustration: Phantom Validation

## Slide 28: Illustration: Phantom Validation

## Slide 29: Illustration: In Vivo Results
>1
<1
Intensity ratio 890 nm / 720 nm
Figure from: Spectrum Medical, Non-Invasive Measurement of SaO2, SvO2 & Hemoglobin

## Slide 30: How deep can we image?
Increasing depth from skin
medial-lateral depth color-coded PAT MIPs
48 mm

## Slide 31: UST: Mathematical Modelling (simplified)
acoustic wave equation
measurement (on boundary)
multiple sources
inverse problem:
acoustic parameter identification problem with boundary data.

## Slide 32: UST Reconstruction Approaches
Travel time tomography: geometrical optics approximation
robust & computationally efficient (time-of-flight picking followed by linear inversion)
valid for high frequencies (attenuation!), low res, lots of data
Javaherian, L, Cox, 2020. Refraction-corrected ray-based inversion for three-dimensional ultrasound tomography of the breast. Inverse Problems
Full waveform inversion (FWI): fit full wave model to all data
high res from little data, transducer modelling, constraints
many wave simulations, complex numerical optimization
lower TRL but already used in 2D systems
our system is not designed for it

## Slide 33: Time Domain Full Waveform Inversion
Pérez-Liva, Herraiz, Udías, Miller, Cox, Treeby, 2017. Time domain reconstruction of sound speed and attenuation in ultrasound computed tomography using full wave inversion. JASA
gradient for first-order optimization via adjoint state method
where is the time-reversed data discrepancy
two wave simulations for one gradient (for one source )
Starting point in 2D:
, s.t.

## Slide 34: 3D Time Domain FWI for Breast UST
Challenges and solutions for 3D:
2 x wave simulations per gradient  stochastic quasi-newton optimization (SL-BFGS)
computationally & stochastically efficient gradient estimator  source encoding for time-invariant systems
memory requirements of gradient computation  time-reversal based gradient computation
slow convergence and local minima  coarse-to-fine multigrid schemes
computational resources  runs on single GPU, can utilize multiple GPUs
,

## Slide 35: Numerical proof-of-concept study
L, Pérez-Liva, Treeby, Cox, 2021. High Resolution 3D Ultrasonic Breast Imaging by Time-Domain Full Waveform Inversion. Inverse Problems

## Slide 36: Numerical proof-of-concept study
L, Pérez-Liva, Treeby, Cox, 2021. High Resolution 3D Ultrasonic Breast Imaging by Time-Domain Full Waveform Inversion. Inverse Problems
Starting point
Improvement

## Slide 37: FWI for Experimental Data: Where Are We?
data from phantom objects, volunteers & patients
ray-based travel time tomography reconstructions
photoacoustic reconstructions  data pre-processing, scanner & transducer modelling, wave simulations
modeling of US protocol, data read-in & pre-processing
model calibration
FWI of phantom objects, quantitative evaluation
FWI of volunteer / patient data
Clinical evaluation

## Slide 38: SOS Phantom
X
Z
SOS – cwater
-100 0 +100
Y
Notes: 50.000 source activations

## Slide 39: SOS Phantom: Travel Time Tomography
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
2.0mm grid; (bent-ray)
SOS – cwater
Notes: ToDo: Redo images, they don‘t seem to be aligned

## Slide 40: SOS Phantom: FWI
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
2.0 mm grid, 0.37/0.21 MHz, 0.22% original signal
SOS – cwater

## Slide 41: SOS Phantom: FWI
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
SOS – cwater
1.5 mm grid, 0.50/0.28 MHz, 0.52% original signal

## Slide 42: SOS Phantom: FWI
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
SOS – cwater
1.0 mm grid, 0.75/0.42 MHz, 2.47% original signal

## Slide 43: SOS Phantom: FWI
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
SOS – cwater
0.8 mm grid, 0.94/0.53 MHz, 6.55% original signal

## Slide 44: SOS Phantom: FWI
X
Y
Z
Y
-100 0 +100
X
Z
SOS – cwater
Y
SOS – cwater
0.6 mm grid, 1.25/0.71 MHz, 21.35% original signal
Notes: This the last result I got. In principle, we can go down to 0.4 mm grids, so I’ll continue a bit. One can see that you get high resolution features, sharp boundaries and a larger field of view compared to the TOF reconstruction. The image contains a few artifacts. Some of them come from the phantom design (reflections from the airsurface, there is another breast cup in there that I did not know about, the interface between the oil and PAA has areas of higher and lower density and gaps in acoustic coupling).

## Slide 45: In Vivo Results: TTT
-100 0 +100
X
Z
SOS – cwater
Y
2.0mm grid
bent ray

## Slide 46: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
2.0mm grid
0.22% original signal
Y

## Slide 47: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
Y
1.5mm grid
0.52% original signal

## Slide 48: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
Y
1.0mm grid
2.47% original signal

## Slide 49: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
Y
0.8mm grid
6.55% original signal

## Slide 50: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
Y
0.6mm grid
21.35% original signal

## Slide 51: In Vivo Results FWI
-100 0 +100
X
Z
SOS – cwater
Y
0.4mm grid
65.46% original signal
PAT shows that these three dots are blood vessels

## Slide 52: Improvement (up to now…)
Y
X
Z
SOS – cwater
Y
X
Z
SOS – cwater
Y
-100 0 +100

## Slide 53
Y
-100 0 +100
X
Z
SOS – cwater
Y
X
Z
SOS – cwater
Y
Improvement (up to now…)
Notes: What I find amazing is that this improvement relies on using a mathematical model of wave-matter interaction with the help a big bag of applied mathematics tricks: From numerical wave propagation, stochastic optimization, multi-grid schemes and other things.

## Slide 54: Using 2SOS model in PAT
Notes: Colour codes depth, everything above 3cm is clipped off

## Slide 55: Using TOF model in PAT
Notes: Colour codes depth, everything above 3cm is clipped off

## Slide 56: Using FWI model in PAT
Notes: Colour codes depth, everything above 3cm is clipped off

## Slide 57: Summary & Outlook
need for novel breast imaging techniques
photoacoustic (PAT) and ultrasound tomography (UST) give access to high-quality images of optical and acoustic tissue properties
combined PAT+UST scanner designed & build
evaluation on data from phantoms, volunteers & patients
PAT reconstructions work well
proof-of-concept studies of FWI for high resolution 3D UST
demonstration of FWI on experimental data
optical inversion in PAT (QPAT)
L, Pérez-Liva, Treeby, Cox, 2021. High Resolution 3D Ultrasonic Breast Imaging by Time-Domain Full Waveform Inversion. Inverse Problems
Felix.Lucka@cwi.nl Photoacoustic and Ultrasonic Tomography for Breast Imaging
Dantuma et al., 2023. Fully three-dimensional sound speed-corrected multi-wavelength photoacoustic breast tomography. To appear in Light: Science and Applications, arXiv:2308.06754
Notes: reconstruction of experimental data